## Slide 1: Numerical investigation of coincidence-correction in gamma spectroscopy
By Bogye Balázs
Supervisor: Horváth Ákos
Goal: simulate HPGE detector get efficiency curve + get data from literature  calculate coincidence correction for given isotope  analyze spectrum of an old experiment  investigate effect of correction

## Slide 2: Coincidence summing
Radioactive decay  nucleus in excited state  deexcitation - not necessarily 1 step, multiple deexcitations  gamma cascades
Fast (ps) transitions to ground (/isomer) state
  Usually gamma emission
  Can be internal conversion (e- instead of gamma), internal pair production (e- + e+)
Multiple transitions detected as one (ps delay << μs detector time resol.)
Coincidences sum up as real transition peak  summing-in
Coincidences sum up  single peaks not detected  summing-out

## Slide 3: Model: Semkow matrix formalism, a bit modified
Goal: coincidence corrected detection chance of j  i transitions
Inputs:
  Transition probability matrix (X=G+E+W, actually 3 matrices) (from literature)
  Efficiency matrices ( μ(E) , total, photopeak) (from simulation)
  Feeding rates (f) (from literature) (distribution of starting levels)
Full, Partial, Null detection probabilities for each transition
   F, P, N detection chains (from E level j to i, through all possible routes)
Summing up as Eji = ft * Ntj * Fji * Ni1
http://atomfizika.elte.hu/akos/orak/mkm/coszu.pdf

## Slide 4: Geant4 simulation
Geant4 - simulation of the passage of particles through matter
Geometry of HPGE detector:
  Ge crystal (technical datasheet)
  Outer Al shell (measured)
  Inner positions, other elements (technical drawing from internet)
Penelope physics (good at low energy too)
Params:
  source distance
  gamma E, angle (mono energy, distribution)
  geometry complexity
  238U decay simulation
  CaCO3 sample simulation
Various validation (/fun) simulations
Simulation geometry with point source
Technical drawing

## Slide 5: Efficiencies – energy dependence

## Slide 6: Efficiencies – distance dependence
r-2 with r0 effective distance (in detector) 
Fit lines to get r0 intercepts
r0 is energy dependent, different at low (<10 mm) and large (10-100 mm) distances
Fitted with empirical formula from literature

## Slide 7
Energy and distance dependent error
Could be better but is good enough at 100 keV – 3 MeV

## Slide 8: 208Tl
Data from NuDat2 (x (Ig), f, alpha)
Good because:
  feeding factor of 2614 keV = 0
  almost all cascades go through 2614 keV
  guaranteed coincidences
Calculated corrected intensities:
Summing-in
  Impossible transitions
  3 strong summing-in effects
Summing-out
Coincidence corrected / uncorrected intensities for 208Tl
To E level [keV]
From E level [keV]

## Slide 9: Measurement
1-day measurement of mineral deposit from thermal pipe from Bük
5-day background measurement
Calculated 208Tl peak areas ( only this done by me)
Spill spectrum analyzer software, spectrum of sample
Multipeak fitting in python (Spill can’t handle it)

## Slide 10: Result
Uncorrected activity: 76.4 ± 5 Bq
Corrected activity: 96 ± 7.6 Bq
  ≈+25%
Last 3 points much better fit
  4 000 - 80 000 Bq  ≈50 Bq
3 Blue dots at 1000s of Bq
8.7 mm was minimal std dev. of activity
(not realistic distance)

## Slide 11: Possible further research
Sample self-absorption (significant at low energies)
Better distance dependence (exact solid angle)
Simulate dead layer of detector
Calculate activity for all isotopes
Simulate readout
Simulation validation with experiment

## Slide 12: Thank you for your attention